## IROC Houston Head Proton Phantom

Guidelines for Planning and Treating the IROC Houston Proton Head Phantom.

Revised July 2013

The irradiation of this phantom is part of the process for approval of the use of spot scanning proton therapy in clinical trials. The purpose of the phantom treatment experiment is to confirm that the dose distribution planned by each institution can be delivered by that institution, and correctly submitted to the IROC Houston.

The IROC Houston QA Center is requesting that each institution keep the phantom for a period of time no more than 2 weeks.  During this two-week period, the institution will image, plan, and treat the phantom and return it to the IROC Houston. Thank you for your cooperation.

This phantom has been designed and constructed by the IROC Houston. IROC Houston phantom contains two inserts. The water fillable imaging insert contains the GTV (tumor target). The dosimetric insert contains TLD at 2 locations and perpendicular sheets of film in order to evaluate the dose to target.

If you have any questions, please contact the appropriate person.

IROC Houston	Nadia Hernandez	(713) 745-8989	   nhernand@mdanderson.org

IROC Houston	Paige Summers	(713) 745-8989    pasummers@mdanderson.org

DOSIMETRY INFORMATION TO BE SUBMITTED:

The following information is to be submitted to the IROC Houston (include it in the shipping box):

- A completed IROC Houston Head Phantom Institution Information form.

- Original hard copy of the plans and isodose distributions in the sagittal and coronal plane through the target center.

The following information is to be submitted to the IROC Houston (results will not be analyzed until the digital treatment plan is received:

Please follow the login URL: https://mdandersonorg.sharefile.com

Login information:

Username: trangnguyen@mdanderson.org

Password:Phantom8989

- Click on folder named IROC Houston head proton phantom,  select the Add Folder tab on the top right hand side of the screen.  In the folder name box, enter your institution name, city and state, then click Create Folder.

  - Select the folder that you have created, then select Upload Files tab on the right hand side. In the Details box please type in phantom type, irradiation date, and physicist name. Follow the instruction and upload your files. Select Send email notification box when done. Lastly Click Upload Files.

The files to export to the IROC Houston are the digital data for your head proton phantom irradiation in DICOM format, and include all CT slices, 3D composite dose file, structure file and plan file.

Please log out once you finish and inform the IROC Houston by email, Nhernand@mdanderson.org .

DOSE PRESCRIPTION:

The total dose to the phantom is 5.4 Gy, with the following constraints:

- The target (GTV) should be uniformly expanded 5 mm to form a PTV

- Total dose of 5.4 Gy to at least 95% of the PTV,

- A minimum dose of 5.1 Gy to the PTV and

- A maximum dose of 6.2 Gy may be given to < 0.03 cc of the PTV.

The phantom should be imaged, planned and irradiated as if it were an actual protocol patient, incorporating all of your customary quality assurance checks.

IRRADIATING THE PHANTOM

Material included in phantom case:

- Head Phantom

- Solid Dosimetric insert

- Hollow Imaging insert

- TLD Block

- Envelope with background film (hidden from your view; please don’t try to find it)

- Mailing label to return case to IROC Houston at the IROC Houston’s cost

Procedures:

- Fill the imaging insert with water:

  - Remove the plug from the top of the insert.

  - Fill slowly with water.  You may need to jiggle the insert to release air trapped inside the cavity.

  - Replace the plug.

2.	Position the imaging insert. The side without plug and hole should be inserted toward the top of the head. Align the black lines on the inferior end of the insert and phantom.  Make sure that the insert is in its correct position by making small rotations of the insert around its central axis. When it is in the correct position it will be locked in place by an indentation at the end of the insert.

- Position and CT the phantom as you would a patient. You may wish to scan with thin slices, especially near the center to better identify the target. Remove the imaging insert. Drain the water from the insert. Place the insert in the box.

- Position the dosimetric insert. The side without a handle should be inserted.  Align the two black lines.  As in step 2, be sure that the insert is in its correct position.

- CT the phantom again with the same scan protocol as the previous scan.

- Fuse the two CT image sets in your treatment planning system.

- Segment the phantom images contouring the target, the TLD, and the outline of the dosimetry insert.

Please see Proton Treatment Planning Addendum for further instruction.

- For treatment, ensure that the dosimetric insert is still in the proper position from the CT scan.

- Perform your customary QA of the proton plan prior to irradiating the phantom. Include in the form values and all the information you consider relevant for analysis.

- Irradiate the IROC Houston TLD block according to the instructions provided.

- Treat the phantom with the developed plan as you would a protocol patient.

- Remove the dosimetric insert and place it in the box.

- Put the empty phantom in the box.

- Include the dosimetry data discussed above. Complete the attached forms. Be sure to include the scale used on the images coming from your TPS.

- Return the complete package to the IROC Houston.

###### IROC Houston Head Phantom Institution Information

Institution:

Address:

Person performing irradiation:

Physicist to receive report:

Oncologist to receive report:

Person to call in case of questions:

Phone Number:	Fax Number:

Email address: _________________________________________________________________

## Treatment Unit:

Manufacturer:	Model:

In-house specification:	Serial #

Proton Energy:   Nom ________(MeV) Range: ______  cm

1. For the phantom irradiation, technique used was (check one)

  IMPT (variable intensity spot scanning).  Continue to next section

  Spot Scanning (energy stacking).  Continue to #4

  Uniform Scanning.  Continue to #2

  Passive Scattering.  Continue to #2

2. Collimation technique:                    	       Multileaf                                 Solid Aperture

3. Range modulation technique:

  Range modulator wheel                    Range shifters

  Both RMW and shifters                         Other, please describe

4.  Compensator technique:

  Solid compensator / bolus                      Other, please describe

Please enclose original copies of your treatment plans.  Include the coronal and saggital planes through the target center. Include scaling factors for each plane. FTP the digital treatment plan.

Treatment Planning System:

Manufacturer: _______________________________   Model: _______________________

Software:________________Algorithm:_______________    Version___________________

Number: ________________________

Treatment of Phantom:

Date of Irradiation:

### Dose specified is to:	 Muscle 	 Water

and is :      	 Physical                                Biological - RBE used is ______

Indicate the dose delivered to these specific points as determined by your treatment planning computer

| Point | Dose (Gy) |
| --- | --- |
| TLD position – Left Posterior |  |
| TLD position – Right Anterior |  |

Results of QA:___________________________________________________________ ______

______________________________________________________________________________

Did you change the M.U. based on your QA? No Yes

Comments:

| For Office Use Only | TLD Batch | Film Batch | Phantom ID # | Code | Date Sent |  | Date Rec'd |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CFA | B11 | A11291201 |  |  |  |  |  |  |

This is a cross-sectional view of the phantom with the dosimetry insert.

Anterior

Sagittal Film Plane

Target

Left

Right

Coronal Film Plane

TLD

Posterior

Notes:

- You need to deliver 5.4 Gy to the PTV (in 1 or more fractions).

- Please ignore all markings on the external shell of the phantom, use your own system to position the phantom.

Proton Treatment Planning Addendum

Please create a plan using the following guidelines:

The total dose to the phantom is 5.4 Gy, with the following constraints:

- Tumor GTV – you may use this structure as the PTV

  - Total dose of 5.4 Gy to at least 95% of the PTV

  - A minimum dose of 5.1 Gy to the PTV and

  - A maximum dose of 6.2 Gy may be given to < 0.03 cc of the PTV.

Once the plan has been created and saved please contour the entire insert highlighted in yellow (see below) and change the relative stopping power to 0.997 to correspond to the dosimetry insert. Using the original optimization, calculate the dose to the dosimetry insert.  This way, we can directly compare our measured dose profiles to a calculated dose profile from the insert containing the dosimeters. Please submit both plans to the IROC Houston.

The yellow overlay depicts the location of the insert for this phantom. The imaging insert contains the anatomy while the separate polyethylene dosimetry insert contains the film and TLD.