INSTRUCTIONS FOR SRS QUALITY AUDIT SYSTEM

Enclosed is a head phantom with two inserts; an imaging insert containing the "tumor," and a dosimetry insert containing film and TLD.  The purpose of the phantom is to verify dose and dose distribution, as well as localization of the dose relative to a simulated “tumor."

The imaging insert contains a nylon sphere suspended in water as the target volume (tumor).  The dosimetry insert contains two orthogonal sheets of GAFchromic film passing through the center of the target and two TLD within 0.5 cm of the center of the target (see the figure).

In order to ensure optimum precision of our evaluation, the following constraints should be observed in planning the treatment:

- The maximum dose to the film should be approximately 30 Gy.

- The dose to the TLD (at the center of the target) should be approximately 30 Gy.

Our report to you will include the following:

- The dose at the center of the target (nylon ball) measured by TLD.

- Dose profiles measured by film.

Image, plan, and treat this phantom as if it were a real patient.  Please send us enough documentation so we can tell what you have done.  Please follow the enclosed instructions and return the phantom and paperwork as soon as possible.

Nylon Ball

Imaging insert

Alignment Pin

Imaging insert

Dosimetry Insert

The phantom has been shipped to you with the following:

- Anthropomorphic Phantom Shell

- Imaging Insert

- Dosimetry Insert

- Drain Plug Tool (blue, white or optional)

- Securing plate

RETURN SHIPPING INSTRUCTIONS

Please return the phantom using the enclosed prepaid mailing label.

For technical questions and questions regarding the return shipment contact

Trang Nguyen	(713)-745-8989,	trangnguyen@mdanderson.org

Nadia Hernandez 	(713)-745-8989, 	Nhernand@mdanderson.org

Paola Alvarez 	(713)-745-8989, 	palvarez@mdanderson.org

Paige Summer 	(713) 745-8989, 	PASummer@mdanderson.org

DOSIMETRY INFORMATION TO BE SUBMITTED

The following information is to be submitted to the IROC Houston (results will not be analyzed until the digital treatment plan is received):

Please follow the login URL: https://mdandersonorg.sharefile.com

Login information:

Username: trangnguyen@mdanderson.org

Password: Phantom8989

- Click on folder named IROC Houston SRS phantom folder; select the Add Folder tab on the top right hand side of the screen.  In the folder name box, enter your institution name, city and state, then click Create Folder.

- Select the folder that you have created, then select Upload Files tab on the right hand side. In the Details box please type in phantom type, irradiation date, and physicist name. Follow the instruction and upload your files. Select Send email notification box when done. Lastly Click Upload Files.

The files to export to the IROC Houston are the digital data for your SRS phantom irradiation in DICOM format, and include all CT slices, 3D composite dose file, structure file and plan file.

Please log out once you finish and inform the IROC Houston by email, Nhernand@mdanderson.org.

Do not send MRI data set , to analyze the phantom irradiation CT data set are used, not MRI.

PROCEDURES

Imaging:

- Remove the large white filler plug on the head phantom, using the enclosed blue tool if necessary, and the small red thumbscrew, which is a vent plug.  Fill the phantom with water, rotating the phantom to minimize the amount of trapped air. It will be necessary to overfill the head to ensure that it is full. Check for excessive air and add more water if necessary. When the phantom is filled, first secure the large white plug, then the small red thumbscrew on the vent hole.

- Remove the small white bolt in the center of the Imaging Insert and fill the insert with water, minimizing the amount of trapped air by overfilling. Replace the bolt; ensuring black ring is in place. Hand tightening should be sufficient.

- Ensure that all bolts and plugs are tight and not leaking. Use the blue tool to tighten the large filler plug if necessary.  Dry off all excess water from the phantom and insert. Be sure to dry the surfaces of the insert receptacle.

- Insert the imaging insert into the neck of the phantom. The flat side, with no writing on it, should be inserted into the phantom first. Press in very slowly, allowing air to escape. The approximate position is obtained by rotating the insert until the yellow stripe on the insert is aligned with the yellow stripe on the head shell.  Make sure the insert is inserted completely into the head cylinder so that there is no gap between the insert and the end of the head cylinder.  Attach the securing plate over the imaging insert using the two green screws provided.

- If you use a headframe, position it on the phantom.  The target is located roughly between the upper parts of the ears.  The dose profiles will be in the sagittal and coronal planes of the phantom; mount the headframe so that these planes correspond with the dose plane produced by your treatment planning system. If possible have your physician participate in the mounting of the frame.

- Locate the target with CT, if preferred you can also use MRI.  If you use both CT and MRI, make sure both set of images are fused.  Please provide axial, coronal and sagittal images through the target volume.

- Remove the two green screws, which hold the securing plate.

- Remove the Imaging Insert. It will come out very slowly due to the tight fit.

Planning:

- The target is the nylon ball that should be evident in either the CT or the MRI scans. The center of the ball defines the center of the target volume. Please treat the nylon ball as the PTV and ignore any air pocket inside the nylon ball.

- Plan a treatment that covers the target volume while sparing all other tissues. To achieve maximum precision, the IROC Houston would like the dose to the center of the target to be approximately 30Gy (100% isodose line). This should be achievable with the following dose guidelines:

  - GammaKnife: Cover the target with 15 Gy (prescription isodose line 40% - 60%)

  - CyberKnife: Cover the target with 20 Gy (prescription isodose line 60% - 80%)

  - C-arm linear accelerator: Cover the target with 25 Gy (prescription isodose line >85%)

- If you cannot meet this guideline (e.g. isodose line is >20% different from the suggested value), please contact the IROC Houston.

Treatment:

- Insert the Dosimetry Insert (white plastic). The dosimetry insert should be placed in the phantom with the white section first so that the writing is visible. Press in very slowly, allowing air to escape. The approximate position is obtained by aligning the thick green lines on the phantom and the insert.

- Attach the securing plate over the imaging insert using the two green screws provided. Notice the red line on one side of the securing plate. There is an alignment pin between the insert and the securing plate; you may have to rotate the insert in order to attach the securing plate.

- Irradiate the phantom according to your plan.

- Remove the dosimetry insert. Lift the insert out first.

- Fill out the data sheet.

- Upload the DICOM data on the IROC Houston sharefile.

- Drain the water from the phantom and the Imaging Insert.

- Wipe excess water off all surfaces of the phantom.

- Pack and return all items. Please include copies of treatment plans and diagnostic images used to plan the treatment.

Stereotactic Radiosurgery Quality Audit Data Form

Institution:

Address:

Person performing irradiation:

Person to receive report:

Person to call in case of questions:

Phone Number: _____________ Email address: _______________Fax Number:

## Treatment Unit (Linear Accelerator or Gamma Knife):

Manufacturer:	Model:

In-house specification:

If linac:

Photon Energy:		Ionization Ratio (TMR 20/TMR 10):

Collimator (if appropriate):

Stereotactic System (if modification to linac):

Manufacturer:	Model:

Please enclose original copies of your treatment plans.

If gamma knife:

Collimators plugged (if any):

Immobilization System (Head frame):

Manufacturer:	Model:

Treatment Planning Syst:

Treatment Planning System: ______________________. Manufacturer:____________

Software:  _________________________ Version Number: _____________________

Treatment of Phantom:

Date of Irradiation: ____________________________

Indicate the dose delivered to the TLD as determined by your treatment planning computer. To know the position of both TLD (superior TLD and inferior TLD) refer to the attachment with the specific coordinates for the phantom you receive.

| TLD | Dose (Gy) |
| --- | --- |
| TLD superior |  |
| TLD inferior |  |

### Dose specified is to:	___ Muscle 	___ Water

(Note:  Several of the following may be the same.)

Maximum dose:		Gy. Isodose line this represents ___________%

Dose to center of the target:_______	Gy. Isodose line this represents___________%

Prescription dose:___________	Gy. Isodose line this represents:__________%.

Minimum dose to target:	______	Gy. Isodose line this represents:	___________%

Diameter of treatment volume (identified by prescription isodose line) along: (Required). Determined from your treatment planing system.

Right/Left Profile:

Anterior/Posterior Profile:

Superior/Inferior Profile:

Comments:

__________________________________________________________________________________

__________________________________________________________________________________

| For Office Use Only | TLD Batch 11 | Phantom ID # | Code | Date Sent |  | Date Rec'd |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | FILM BATCH S0628 MD-V2-55 |  |  |  |  |  |  |