Reference Cases

Please calculate and attach isodose distributions for the single seed and geometric implant described below.  Sources should be the model 125I or 103Pd seeds (from the list of models complying with the AAPM prerequisites – see http://rpc.mdanderson.org) that you choose to treat your patients, with source strength specified exact at the beginning of the implant.  Do the calculations as you would do them clinically using the TG-43 dosimetry, detailing any assumptions necessary.

Case 1:	A single seed, strength 0.5U for I-125 (Gy m2 h-1) (use 2.5 U for 103Pd; use 1.8 U for Cs-131). If your software allows and you use a line source approximation, calculate both in the longitudinal and mid-transverse planes of the seed.  Please submit isodose lines from 0.2 to 100 Gy.  (Lines 0.2, 0.5, 1, 5, 10, 50, & 100 Gy are preferred.) The lower isodose lines can be achieved by using a lower prescription dose.

Dosimetry Calculations:

Write below the equation that will be used for hand calculating the instantaneous dose-rate to an arbitrary point from a single seed in the TG-43 formalism.  (If possible give notations used by your treatment planning computer).  The intent is for you to be able to verify that the values of various parameters in your treatment planning system are the same as in TG-43.

Define the variables in the equation:

For each seed model used to treat patients on this protocol, submit the data used by your treatment planning computer for the following parameters:

- Dose rate constant Λ

- Anisotropy constant φ and/or factors

- Radial dose function

- The units of SK are :     ___________________

- Do your 125I dose calculations agree with TG-43 to within +5% from 5-70 mm? Yes     No

- Do your 103Pd dose calculations agree with TG-43 to within +5% from 5-50 mm? Yes     No

Note:	If you wish to use more than one of the approved seed models, please submit case 1 for each seed model.