8060 El Rio St.  Houston, TX, 77054   	● IROCHouston@mdanderson.org

Tel (713) 745-8989 	● Fax (713) 794-1364  	● http://IROCHouston.mdanderson.org

Craniospinal Irradiation Benchmark

This benchmark is a sample case used to evaluate the treatment planning process for craniospinal irradiation (CSI) at your institution, i.e., data acquisition, treatment planning, dose calculation and monitor unit calculations. The aim is to demonstrate your capability to participate fully in protocol studies requiring CSI.  Your benchmark will be evaluated by the IROC Houston QA Center to assess the accuracy of your dose calculation, and will also be shared with QARC (Quality Assurance Review Center). It will cover all protocols requiring CSI; you will not be expected to complete separate benchmarks for different CSI protocols.

The CT image set may be downloaded from the IROC Houston website

(http://IROCHouston.mdanderson.org) or it can be sent to you on CD (request by e-mail from IROCHouston@mdanderson.org).

Data should be submitted by hardcopy, two (2) full sets of data should be sent to the address below.

Please return completed forms and supporting documents to:

UT MD Anderson Cancer Center

IROC Houston QA Center

c/o Dosimetry

8060 El Rio St.

Houston, TX 77054

Phone:  (713) 745-8989

Fax:  (713) 794-1364

Email:  IROCHouston@mdanderson.org

Section 1.  Description of Craniospinal Irradiation (CSI)  Benchmark

Purpose:

This benchmark is a sample case used to evaluate the treatment planning process for CSI irradiation at your institution, i.e., data acquisition, treatment planning, dose calculation and monitor unit calculations. The aim is to demonstrate your capability to participate fully in protocol studies requiring CSI.  This benchmark will be used by both QARC (Quality Assurance Review Center) and the IROC Houston QA Center to assess your treatment planning and your dose calculation algorithm.  It will cover all protocols requiring CSI; you will not be expected to complete separate benchmarks for different CSI protocols.

1. Method:

A CT scan set in DICOM format is to be loaded into the treatment planning system that is used for planning protocol patients.

2.  Treatment Prescription:

The prescription dose for each treatment site (brain and spinal cord) is 24 Gy in 1.5 Gy fractions.

You should plan this treatment as is the normal practice at your institution, that is either CT-based or not.

A)  CT based:  Plan the opposed lateral brain fields and the spinal irradiation as you normally do.  Pay particular attention to the junction of the brain and spine fields.

B)  2D, non-CT based:  Design the brain fields using digitally reconstructed radiographs (DRRs) from the CT scan. Perform the dose calculation as you normally would. Determine the size of the spinal irradiation field from the CT.  Determine any couch or collimator rotation as you normally do in your clinic.  If you routinely calculate a “gap” between the brain fields and spinal field, do so.  Perform the dose calculation as you normally would.

Section 2.  Data to be Submitted

Two (2) full sets of the following shall be submitted as original hardcopy and in color. Please make sure that isodose contours are readily identifiable and that a scale is included on all images and Beam’s Eye Views.

For CT-based plans:

1. Axial CT with isodose contours through the isocenter of the brain fields.

2. Sagittal section through the spinal cord with isodose contours.

3. DRR in the Beam’s Eye View (BEV) for each field, showing the aperture.

4. A printout of beam specifications including at a minimum the beam energy, gantry, couch, and collimator positions, field sizes, apertures, wedge specifications, depth of isocenter (or SSD) and monitor unit calculations.

5. A description of your method of abutting the cranial and spinal fields. Include all forms and calculations you employ. If you usually use a “gap”, include the complete gap calculation as performed in your clinic.

6. Complete and submit the CSI questionnaire below.

7. Complete and submit the Dosimetry Summary Form.

For non- CT-based (2D) plans:

1. Printout of the shape of each of the treatment fields.

2. A printout of beam specifications including at a minimum the beam energy, gantry, couch, and collimator positions, field sizes, aperture names, wedge specifications, depth of isocenter (or SSD) and monitor unit calculations.

3. A description of your method of abutting the cranial and spinal fields. Include all forms and calculations you employ. If you usually use a “gap”, include the complete gap calculation as performed in your clinic.

4. Complete and submit the CSI questionnaire below.

5. Complete and submit the Dosimetry Summary Form.

Craniospinal Irradiation

Benchmark Questionnaire

Institution:  _______________________________  City ____________ State________

Treatment Planner: ___________________________   Telephone:________________

e-mail: __________________________________          Fax:  _____________________

Brain Fields:

Beam energy:	___________ MV

Collimator rotation (if any):  ________o

Couch Rotation (if any):  ___________o

Gantry Rotation (if any):  ___________o

Monitor Units: Field 1: ____________ MU

Field 2: ____________ MU

Spine Field:

Beam energy/modality:	______ MV or _______ MeV (electrons)

SSD     or   SAD    Setup

Collimator rotation (if any):  ________o

Couch Rotation (if any):  ___________o

Gantry Rotation (if any):  ___________o

Monitor Units: ____________ MU

Abutment of the Cranial and Spinal Fields:

 Abut light fields at posterior neck skin

 Calculate “gap”

 Always use ___ cm “gap”

 Other (describe) ________________________________

Is the matchline shifted (feathered) routinely? Yes       No 

If yes, at what interval (e.g. every 5 fractions): __________________

and by how much (e.g. 1 cm): ___________

RADIOLOGICAL PHYSICS CENTER

DOSIMETRY SUMMARY FORM

INSTITUTION NAME:

PHYSICIST/DOSIMETRIST:

#### DOSE PRESCRIPTION

| Number of Fractions |  | Dose to Prescription Point | cGy | Treatment Machine |  |
| --- | --- | --- | --- | --- | --- |

#### TREATMENT TECHNIQUE

|  | FIELD I | FIELD II | FIELD III | FIELD IV | FIELD V | FIELD VI |
| --- | --- | --- | --- | --- | --- | --- |
| FIELD NAME (ANT, POST, RT LAT, etc.) |  |  |  |  |  |  |
| ENERGY, MODALITY (e.g. 6X, 6e) |  |  |  |  |  |  |
| SSD |  |  |  |  |  |  |
| GANTRY / COLLIMATOR / COUCH ANGLE | / | / | / | / | / | / |
| COLLIMATOR SETTING (W x L or if using independent jaws X1, X2, Y1, Y2) | x X1 X2 Y1 Y2 | x X1 X2 Y1 Y2 | x X1 X2 Y1 Y2 | x X1 X2 Y1 Y2 | x X1 X2 Y1 Y2 | x X1 X2 Y1 Y2 |
| EFFECTIVE BLOCKED FIELD SIZE |  |  |  |  |  |  |
| DEPTH OF PRESCRIPTION POINT |  |  |  |  |  |  |
| OFF-AXIS (DISTANCE/FACTOR) | cm/ | cm/ | cm/ | cm/ | cm/ | cm/ |
| ATTENUATORS: Blocks or Multi-leaf collimators (MLC) (tray factor if applicable) |  |  |  |  |  |  |
| Wedge (angle/factor) ‪ Physical ‪ Internal ‪ Enhanced Dynamic ‪ Virtual | / | / | / | / | / | / |
| Compensator (factor, material, thickness) / Bolus (depth) |  |  |  |  |  |  |
| DOSE PER FRACTION |  |  |  |  |  |  |
| MONITOR UNITS PER FRACTION |  |  |  |  |  |  |