TRUS Verification Quality Assurance

Institution:____________________________________ Date: _________________________

Name of the person performing the QA: __________________________________________

Ultrasound unit	Manufacturer:___________ 	Model:____________	S/N:_____________

Probe			Frequency (MHZ) :_______	Model:____________	S/N:_____________

Phantom 		Manufacturer:___________ 	Model:____________	S/N:_____________

- Grayscale Visibility Test:

Results: (Action limit: change>2 steps or 10% of baseline value)

| Measured number of discreet steps/Gradient Distance: |  |
| --- | --- |
| Baseline Value: |  |
| Agreement: |  |

- Depth of Penetration Test:

Results: (Action limit: change > 1cm from baseline value)

|  | Axial Plane | Longitudinal Plane |
| --- | --- | --- |
| Depth at which speckle pattern can be separated from electronic noise: |  |  |
| Baseline Value: |  |  |
| Agreement: |  |  |

- Axial and Lateral Resolution Test

Results: (Action limit: change >1 mm from baseline value)

|  | Axial Plane |  |  | Longitudinal Plan |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Current | Baseline | Difference (Abs/%) | Current | Baseline | Difference (Abs/%) |
| Axial Resolution: |  |  |  |  |  |  |
| Proximal |  |  |  |  |  |  |
| Distal |  |  |  |  |  |  |
| Lateral Resolution: |  |  |  |  |  |  |
| Proximal |  |  |  |  |  |  |
| Distal |  |  |  |  |  |  |

- Axial and Lateral Distance Measurement Accuracy Test:

Results: (action level: Axial – difference >2 mm or 2% from nominal value; Lateral – difference >3mm or 3% from nominal value)

|  | Axial | Lateral (proximal) | Lateral (distal) | Sketch: |
| --- | --- | --- | --- | --- |
| Expected distance |  |  |  |  |
| Measured distance |  |  |  |  |
| Absolute difference |  |  |  |  |
| Percent Difference |  |  |  |  |

PAGE 2: TRUS Verification Quality Assurance

- Distance, Area and Volume Measurement Accuracy Test:

Results: (Action level: Difference >2 mm or 2% for distance; difference >5% of nominal for area and volume measurements)

|  | Top-Bottom Distance | Left-Right Distance | Area Measured | Volume Measured |
| --- | --- | --- | --- | --- |
| Expected |  |  |  |  |
| Measured |  |  |  |  |
| Absolute Difference |  |  |  |  |
| Percent Difference |  |  |  |  |

- Needle Template/Electronic Grid Alignment Test:

Results: (Action level: alignment should be correct to within 3mm)

| Coordinate of Greatest Discrepancy | Error (mm) |
| --- | --- |

- Treatment Planning System Volume Test

Results: (Action limit: The volumes calculated by the ultrasound system and the treatment planning computer should agree to within 5%)

|  | Target Volume (cc) |  | Absolute difference (cc) | Percent Difference (%) |
| --- | --- | --- | --- | --- |
| Known volume |  |  |  |  |
| Measured with ultrasound |  | Ultrasound volume relative to known volume |  |  |
| Measured with treatment planning system |  | TPS volume relative to ultrasound volume |  |  |

Notes:

- It is reasonable to coordinate testing with the vendors, as long as you are confident that the vendor is competent in the tests he/she is doing.

- It is advised to use the CIRS Model 045 Brachytherapy QA Phantom to follow TG-128, although any phantom that allows the recommended tests is acceptable.

- Axial resolution is essentially independent of depth, however one may measure it at different depths, but it is generally unnecessary.

- All tests should be performed with the probe frequency most commonly used during implants and the maximum power output.