IROC HOUSTON QA CENTER

### COBALT –60 MACHINE DATA

- INSTITUTION:		MACHINE:

- MANUFACTURER:  		MODEL:

- DATE MACHINE STARTED CLINICAL USE:  __/__/__

- DATE PRESENT SOURCE INSTALLED: __/__/__

- SOURCE DIAMETER:  ___ cm  	NOMINAL TREATMENT DISTANCE: ___ cm

- OUTPUT DETERMINATION:

Present protocol:  	  	date implemented:  __/__/__  previous protocol:

- SETUP:  ___ cm x ___ cm, ___ cm S___D, trimmers ___ cm

## Ionization measurements:		   in air		   in phantom

Gantry angle:  	

# Source to Chamber Distance:  ___ cm

# Distance from source to point of stated dose rate:  ___ cm

# phantom:  composition: 	,	density:  	g/cm3

ionization chamber:  	 	thimble (wall) thickness:

   greater

   less

- Monitor end effect:  	 mu.  	Net mu 		than mu set.

used during:    output calibration	  patient treatment

- FACTORS USED TO CALCULATE ABSORBED DOSE RATE (Gy/mu) FROM DOSIMETER READING (give equation, define all factors and give numerical values; attach extra sheet if necessary.  Attach detailed calculation from most recent annual calibration.)

COBALT-60 MACHINE DATA (cont'd)

3.	DOSE PRESCRIPTION INFORMATION:

- Output dose rate is stated to:

# A.		muscle or	 water

# B.		miniphantom	 dmax	 other/explain

# C.	at ____ cm	 SAD	 SSD.

- For protocol patients:  Calculated to muscle (tissue)?

yes     no 	If yes, indicate when this is done.

4.	QUALITY ASSURANCE INFORMATION:

- How often is a full calibration done?	 monthly	 semi-annually	 yearly

- System used for full calibration?	 primary	 secondary

- How often are outputs checked?	 daily	 weekly

other/explain

- How often is dose rate updated for decay?

- Date that dose rate is exact:	  first of the month
		  middle of the month
		  other

- Distance from isocenter to reference point on machine:
	Reference point:  					
	Distance to isocenter:  ________ cm

	How is this distance determined?

- How is treatment distance determined for patients?

using ODI 	using lasers      other, describe

- How often are ODI and lasers compared to mechanical indicator?

- For treatments at distances other than the nominal distance, how is the dose rate 	determined?

		  Inverse square correction	 Measurement

- Treatment techniques used?  	SSD:  _____%	       SAD:  _____%

- Source of depth dose data (published or measured):

Published	Reference